## Slide 1
EECS 289 – 2024 Lab 1
Neurotransmitter sensors and IrOx pH electrodes

## Slide 2
Schematic for electrode fabrication process
Orientation of pads in the electrode
CAD image of the electrode with dimensions
(unit = μm)

## Slide 3
Sensor Fabrication
Higher Sensitivity
Specificity
Selectivity

## Slide 4
Sensor Fabrication
Ascorbic Acid (C6H8O6)
Dopamine (C8H11NO2)
Selectively lets ions or molecules through the aromatic diamine polymer film through size exclusion
Applying different process result in different pore size of the formed polymer
Electropolymerization of mPD
Other electroactive molecules can generate current and interfere with the signal
Exclusion layer to block the molecules from reaching the electrode surface needed

## Slide 5
Self-referencing: Monitors any other signal except for the one from the target
Detects all signals from other electro-active chemicals except for the target
Receives same ambient noise that the sensing site does
Can later be subtracted from sensing site to get higher signal-to-noise ratio (SNR)
Orientation of pads in the electrode
Self-referencing

## Slide 7
Calibration
Potentiostat Pinnacle 3104
+0.7 V bias between WE and RE
30 minutes required to reach the baseline
Calibration
  20 μM of L-glutamate in steps
Interferences
  10 μM of Dopamine – DA
  200 μM of Ascorbic Acid – AA

## Slide 8
Sensor Performance - Sensitivity
Calibration curve and the sensitivity for Pt-black on Pt to sequential addition of L-glutamate (n=4)
Limit of Detection (LOD) = 1.92 μM
R2 = 0.98

## Slide 9
Sensor Performance - Sensitivity
Calibration curve and the sensitivity for Pt-black on Pt to sequential addition of L-glutamate (n=4)
Limit of Detection (LOD) = 1.92 μM
R2 = 0.98

## Slide 10
IrOx electrode as pseudo-RE
Regular RE
Why IrOx?
IrOx is used to sense pH. In the brain, pH slightly changes, so an IrOx electrode can be used as a pseudo-RE
IrOx is biocompatible (see slides in the Neuro-lecture)
Au
AgCl
IrOx
Day I
Day III

## Slide 11
Lab 1 to do list
IrOx electrodes for pH measurement
  pH calibration
  Plot the curve and calculate the sensitivity
Neurotransmitter sensors
  Either L-glutamate sensor or DA sensor with glass RE and IrOx pseudo-RE
  Plot the curves
  Calculate the sensitivity and the LOD